TREND (through LAN) / Sample / Software 				      V 1.1

Advantest Corporation

- Install sample software into PC  (RUN Package\Seutp.exe)

The original version of the installer was for the Japanese version of Windows 2000. You may see strange characters, but proceed as you would for any installation and press yes.

- Set U3700 Series IP address

System – Remote Control – LAN IP address – Network Setting

Fig.1 shows the settings for a Point-to-Point example

If you wish to connect through a LAN network,

please ask your network manager for the IP address.

- Set PC IP address

Required for direct connection. Not necessary for LAN network connection.

For WINDOW 2000

My network – Right click - Properties –

Local Area Connection - Right click - Properties –

Internet Protocol (TCP/IP) – Properties

Set IP Address

This IP address must be set different than the U3700 Series

- Connect LAN cable

You will need a cross over Ethernet cable for a Point-to-Point connection

- Run the Sample software.

START – PROGRAM – U3700 LAN Power Trend － U3700 Power Trend

Set IP address of Spectrum Analyzer（Port No. is 5025）

- Push Connect

When the connection has been established, Connected will be displayed on message line located at the bottom part of window. If you have checked Monitor, the communication log will be displayed.

The parameters of the Spectrum analyzer can be controlled manually using the “Sending Text” command line. Type a GP-IB command in “Sending Text” and push “Send Text” key. The U3751 should respond accordingly.

Send a query command (e.g. CF?) and you can acquire data from the instrument that will appear in the monitor window. If you press SendText with the above mode（＊RST）, Spectrum Analyzer will be initialized.

Before executing Measurement (Power Trend), uncheck the Monitor.

- Measurement (Power Trend)

Press Measurement (Power Trend) and Trend measurement window will be displayed.

Measurement procedure

A. Set up spectrum analyzer parameters.

B. Press Setup SPA（parameters will be sent to Spectrum Analyzer）

(If you would like to set parameters not available from the panel, set the Spectrum Analyzer to local and set up manually.）

C. Press START

D. Measurement will start and the measured data will be displayed.

E. When measurement time has finished, data is saved and the termination massage will be displayed.

Trend measurement screen

○Set up: Center, Span, Reference level and Offset.

○Select Units (common in MKR1 etc)

○Select measurement mode (See below)

○When Fixed or Zone Marker are selected, input measurement frequency.

○Input total measurement time and measurement interval. (The measurement time has some margin of error)

○current setting parameters are saved on Close. (They are automatically loaded at start up)

○You can add a comment. It will be saved with the measurement if the box is checked

- Measurement Modes

A. Fixed Marker	: acquires at specified frequency level (3 points)

B. Zone Search	: search the１Div range of specified frequency (3 points)

C. Peak Search	: search the whole display and finds the result of the largest point.                          （3 points: specifying F is not required）

Note: Max-Min can be used for ABC. The data is the Max data in the Time Interval measurement spacing. Only MKR1 can be get Min data

D. CH Power	: find the result of Channel power. Data is 1 point.

(Channel width is input to CHSP）

- Save measurement data

At the end of measurement, the file named AdvanTrend+MM+HH+NN.dat will be created on the C: drive if Save is checked. (MM is month, HH is time, NN is minute)

If you would like to save it to a different directory, first, you need to create the directory in Windows. Then type the path and file name in the file box - C: \directory\filename. The file name + MM + HH + NN.dat will be created in the directory.

- Measurement Example

- Exiting Program

Complete measurement with CLOSE from measurement window.

Terminate LAN connection using CLOSE from initial screen.