## Slide 1
High Speed Power Measurement
Outline:
This is a sample software where RF signal parameters are measured with external PC by using U3800 Time Domain Analysis (Digitizer).
As a result, 7 or 8 times the speed measurement can be expected compared with standard SPA. (SPA: Spectrum Analyzer)
Measurement items:
The measurement items refer the high performance small electric power data communications system of TELEC. (Telecom Engineering Center) (Note: No Spurious, No EVM measurement)
Measurement method:
Set carrier frequency and band width of the power measurement. Other measurements are measured with SPAN(CB) 1MHz.
*The measurement signal makes 11b_cck_5.5Mbps an example.
*When the pulse width is 1ms or less in other standards, we might need to change the software.
*The measuring range can correspond from special small systems to WLAN.
*The sample software is only an EXE file usually.
*There is an offered service of source file (VB6) by the user registration.
Dec.05.2011

## Slide 2
1. Panel image and Main key
Measure
Stop
is pushed, the measurement begins. The continuous measurement is also possible by check.
is pushed, the measurement is ended. LAN is a waiting state like the opening.
Connect
is pushed, it connects to SPA. Other keys are the prohibitions until the connection is completed.
Local
is pushed, the connection with SPA is released. SPA becomes the local mode.
Input
Results
Show
1-
2-
3-
4-
1-
2-
3-
4-

## Slide 3
2. Measurement method and Detail
CH Power:
Ave Power:
(BW 1MHz)
OBW 99%:
SBW 90%:
Frequency:
After SBW is calculated, the center frequency is calculated. When PS is set, it searches for the peak from the waveform and the frequency is calculated.
DGT and FFT mode, set C-Freq. CB uses the value about twice CHP-BW. After other conditions are set and Cap B and Cap T are displayed, and measuring. CH Power is calculated on specified CHP-BW from the measured waveform.
After frequency is calculated, it converts to P-T graph, and the Average Power is calculated. (In this example, CT is 5ms and AVG Power is 1ms.)
SBW is calculated from the waveform of FFT at CB 1MHz and CT 5ms.
OBW is calculated from the waveform of FFT mode of CH Power.
Startup and Measurement:
a. Execution of the Software. (Select the U3800 PowerMeasDGT from the START)
b. IP Address is confirmed or it changes. (Default:192.168.0.1)
c. When the Connect button is pushed, it is connected with SPA. (Set Using CH, Preset)
d. BW of CH Power is confirmed or is set. (CB and CT are the automatic settings)
e. Select the Trigger Mode. Averaging On/Off is possible while measuring it.
f. Additionally, when a necessary condition is set, and the Measure button is pushed,
the measurement is begun.
g. The end of program pushes the Local, and shuts the Window.
Measurement details of each items:
ACP is calculated from the SBW measurement data if ACP measurement ON. (C-SP and C-BW are input.)
ACP: Low/Up
(Span 1MHz)

## Slide 4
3. Measurement Flow
CH Power:
OBW 99%:
SBW 90%:
Frequency:
Ave Power:
Start
End
Note1: Trigger: Free Run actually uses SI. Sweep End is detected while seeing STB.
Note2: GHz, MHz, kHz, and Hz use a fixed unit for the unit of the display.
Note3: The continuous measurement is enabled.
Note4: Width and Span are up to Max 40MHz.
Setup: IP,CF,SP,REF,
DGT,CB,CT…
Loop:
CHP Meas
SBW Meas
P - T
Get Wide BW data
Get Narrow BW data
FFT to P-T Graph
End?
TO Loop
N
Y
Calculation
Calculation
Calculation
ACP: On/Off
Calculation
Calculation
Calculation

## Slide 5
4. Power Measurement Formula
Formula of Channel Power
(Details: See U3800 series
User’s Guide)
Note: Advantest uses
PBW=RBW*1.06
Formula of Average Power
(Details: See U3800 series
User’s Guide)

## Slide 6
5. Measurement item and Module:1
Example Signal: 11b_cck_5.5Mbps
1. Channel Power (Total Power) BW 22MHz
2. OBW 99%
Measurement１: DGT CB:40MHz, CT:0.3ms, Data(25us):FFT
GetCHPpower CHPstart, CHPstop, MeasCHPCB, CHpower
GetOBWfreq 99, MeasCHPCB, OBWfreq, CentFreq
1.
2.
P-T
(>200u)
Conv.
FFT
(25u)
1.CHpower
2.OBWfreq
(Cent Freq)

## Slide 7
Measurement2: DGT CB:1MHz, CT:5ms, Data(1ms):FFT
3. SBW 90%
4. Frequency
5. AVG Power (1MHz RBW, 1ms)
GetOBWfreq 90, SBWSpan, OBWfreq, CentFreq
Convert the data FFT(Level-F) to P-T(Level-Time)
GetAVGpower AVGpower
3.
4.
5.
Conv.
P-T
P-T
(>5m)
Conv.
FFT(1m)
OBW freq (3.SBW)
Cent Freq (4.Freq)
5.AVGpower
6. Measurement item and Module:2

## Slide 8
7. Change the measuring conditions
Example Signal: 11b_cck_5.5Mbps, Other: See table

## Slide 9
8. Averaging Function
1. Channel Power (Total Power), 2. OBW 99%
>AVG-n (Sweep) is checked, it becomes AVG ON.
The AVG number puts the numerical value in the input column.
It is sweep method. (sweep: 2 to 30 times)
>When the check is put in FFT AVG Mode, AVG is done from one
time base measurement. More FFT is done from one time base
measurement and AVG is done. (FFT: 2 to 10 times)
3. SBW, 4. Frequency
>AVG-m is checked, it becomes AVG ON.
The AVG number puts the numerical value in the input column.
It is sweep method.
5. AVG Power(dBm/1MHz)
>AVG-m is checked, it becomes AVG ON. It is sweep method.
The AVG number uses the same numerical value as SBW.
Averaging is added to each measurement

## Slide 10
9. Detail of SBW, AVG-Power
1. SBW
>When Wide SP is checked, the measuring span of SBW is
widely changed. After OBW99% is calculated from wide span
used when Channel Power is measured,
SBW (90%) is calculated.
2. AVG Power(dBm/1MHz)
>If Peak S is checked, it becomes Peak Search ON.
When Channel Power is measured, it searches for the utmost
level from Freq axial waveform. And, the time base data is
measured at that frequency and mean power calculated.
Detail Functions are added to SWB, AVG-Power